## Stent Tracker
Stent ID | Stent Name | Date Created | Notes | French (Fr) | Length (mm) | OD (mm) | ID (mm) | Wall (mm) | Hole Dia (mm) | Coil R (mm) | Coil L (mm) | Pigtail (°) | Prox Sect (mm) | Dist Sect (mm) | Mid Sect (mm) | r_t | r_sh | r_end | N Prox | N Mid | N Dist | End Hole (mm) | Valid | Unroof (mm) | Run ID | Status | Run Date | Time (min) | Q_out | ΔP | Q_SH1 | Q_SHmid | Q_SHend | Eff % | Notes
STN-001 | Baseline_v1 | 2026-01-20 | Initial baseline design for testing | 6 | 200 | 1.998 | 1.498 | 0.25 | 1.5 | 10 | 30 | 300 | 30 | 30 | 80 | 0.15 | 0.5 | 0.7 | 3 | 5 | 3 | 1.05 | Legacy | 0 | RUN-001 | Completed | 2026-01-20 | 45 | 2.5 | 1200 | 0.8 | 0.6 | 0.7 | 85 | Baseline run successful

## Parameter Reference
Parameter | Description | Unit | Min Value | Max Value | Default | Type
stent_french | Stent outer diameter | French (Fr) | 4 | 8 | 6 | Sampled
stent_length | Total stent length | mm | 100 | 300 | 200 | Sampled
coil_radius | Pigtail coil radius | mm | 5 | 15 | 10 | Sampled
coil_length | Pigtail coil length | mm | 10 | 50 | 30 | Sampled
pigtail_angle | Pigtail wrap angle | degrees | 270 | 360 | 300 | Sampled
section_length_proximal | Proximal section length | mm | 20 | 60 | 30 | Sampled
section_length_distal | Distal section length | mm | 20 | 60 | 30 | Sampled
n_proximal | Number of proximal side holes | count | 0 | 10 | 3 | Sampled
n_middle | Number of middle side holes | count | 0 | 15 | 5 | Sampled
n_distal | Number of distal side holes | count | 0 | 10 | 3 | Sampled
unroofed_length | Unroofed section length at distal end | mm | 0 | 30 | 0 | Sampled
r_t | Wall thickness fraction (t = r_t × OD) | ratio | 0.1 | 0.22 | 0.15 | Sampled
r_sh | Side hole diameter fraction (d_sh = r_sh × ID) | ratio | 0.2 | 0.7 | 0.5 | Sampled
r_end | End hole diameter fraction (d_end = r_end × ID) | ratio | 0.3 | 0.95 | 0.7 | Sampled
OD | Outer diameter = 0.333 × stent_french | mm | — | — | — | Derived
wall_thickness | Wall thickness = r_t × OD | mm | — | — | — | Derived
ID | Inner diameter = OD - 2×wall_thickness | mm | — | — | — | Derived
side_hole_diameter | Side hole diameter = r_sh × ID (capped at ID - 0.15) | mm | — | — | — | Derived
end_hole_diameter | End hole diameter = r_end × ID (capped at ID - 0.15) | mm | — | — | — | Derived
section_length_middle | Middle section = stent_length - 2×coil_length - prox - dist | mm | — | — | — | Derived
VALIDATION CONSTRAINTS
Constraint | Description | Unit | Value | Rule
ID_min | Minimum inner diameter | mm | 0.6 | Reject if ID < ID_min
gap_min | Minimum hole spacing gap | mm | 0.3 | pitch_min = d_sh + gap_min
buffer | End buffer for hole placement | mm | max(1.0, d_sh) | L_use = L_section - 2×buffer
section_middle_min | Minimum middle section length | mm | 10 | Reject if section_middle < 10
d_sh_min | Minimum side hole diameter | mm | 0.3 | Reject if d_sh < 0.3
cap_margin | Hole diameter cap margin | mm | 0.15 | d_sh ≤ ID - cap_margin

## COMSOL Metrics Reference
Metric | Description | Unit | Optimization Goal
Q_out | Total outlet flow rate | mL/min | Maximize
delta_P | Pressure drop (inlet - outlet) | Pa | Minimize
Q_SH1 | Flow through first side hole | mL/min | Monitor
Q_SHmid | Flow through middle side hole | mL/min | Monitor
Q_SHend | Flow through last side hole | mL/min | Monitor

## Run Log
Run ID | Stent ID | Date/Time | Status | Duration (min) | Error Messages | Notes