## Slide 1
Stent Optimization Pipeline
A Walkthrough of the End-to-End Framework
From parametric CAD → CFD simulation → GP surrogate → Bayesian optimization
Akash | Masters Thesis | February 2026
1/18

## Slide 2
Roadmap
① The Problem — why stent design needs a smarter approach
② Pipeline Architecture — the 30,000-foot view
③ Parametric CAD — scripted stent geometry (build123d)
④ Design Space & Sampling — Latin Hypercube in 16 dimensions
⑤ Feasibility Filtering — throwing out impossible designs early
⑥ COMSOL CFD — the expensive simulations
⑦ COMSOL Results Parsing — extracting the data we care about
⑧ GP Surrogate Model — learning to predict without simulating
⑨ Literature Backing — Maulik & Taira (2020)
⑩ Bayesian Optimization — closing the loop
⑪ Orchestration — the full campaign script
⑫ What's Next & Q+A
2/18

## Slide 3
The Problem
Why We Can't Just Guess
• Pediatric ureteral stents drain fluid from the kidney to the bladder
• Current designs are scaled-down adult stents — not optimized for children
• Each stent has ~16 design knobs: diameter, length, hole sizes, hole counts, coil geometry…
• Changing one parameter affects flow, pressure, and structural integrity simultaneously
• Testing every combo physically? Impossible — the design space is enormous
💡 We need a system that explores this space intelligently,
runs simulations only where it matters, and finds the best designs.
3/18

## Slide 4
Pipeline Architecture
The 30,000-Foot View
Design Space (16 dims)
LHS Sampling (space-fill)
Feasibility Filter (constraints)
CAD Generation (build123d)
COMSOL CFD (simulate)
GP Surrogate (learn)
Bayesian Optimization (suggest)
→ → → → → → → → → → → → → → → → → → → → (iterative loop back from BO to COMSOL)
Key idea: We start with a big batch of random designs (LHS), filter the ones that are physically impossible, simulate the rest in COMSOL, then train a cheap statistical model (GP) that learns the patterns. Once the GP is decent, we use Bayesian Optimization to suggest the NEXT best designs to simulate — instead of guessing randomly. This loop repeats until we converge on optimal stent geometries.
4/18

## Slide 5
Project Structure
How the code is organized
masters/ ├── config/parameters.yaml ← all 16 design knobs defined here ├── src/ │ ├── cad/stent_generator.py ← builds 3D stent geometry │ ├── sampling/ │ │ ├── lhs_generator.py ← Latin Hypercube sampling │ │ └── feasibility.py ← rejects bad designs │ ├── comsol/ │ │ ├── runner.py ← fires off COMSOL sims │ │ └── result_parser.py ← reads simulation outputs │ ├── surrogate/ │ │ ├── gp_model.py ← Gaussian Process wrapper │ │ └── training.py ← train + cross-validate │ └── optimization/ │ └── optimizer.py ← Bayesian Optimization loop ├── scripts/ │ ├── run_lhs_batch.py ← CLI: generate a sampling batch │ └── run_optimization_campaign.py ← CLI: full campaign └── docs/ ← decision logs, schemas, metrics
5/18

## Slide 6
The Design Space
16 parameters defined in parameters.yaml
What we sample (the "knobs" we turn):
• stent_french (4–8 Fr) — diameter in French units
• stent_length (100–300 mm) — total body length
• coil_R_prox / coil_R_dist (4–12 mm) — coil radii
• pitch_prox / pitch_dist (4–10 mm) — helix pitch
• turns_prox / turns_dist (0.5–2.5) — number of coil turns
• r_t (0.10–0.22) — wall thickness as fraction of OD
• r_sh (0.20–0.70) — side hole size as fraction of ID
• r_end (0.30–0.95) — end hole size fraction
• n_prox, n_mid, n_dist — number of holes per section
• section_length_prox / dist (20–60 mm)
• unroofed_length (0–30 mm) — half-pipe cut at distal end
Why fractions instead of absolute values? By sampling r_t, r_sh, r_end as ratios (0–1), we guarantee the geometry makes sense by construction. Wall thickness can't exceed the tube radius. Holes can't be bigger than the inner diameter. → Fewer invalid designs = less wasted compute.
Derived (computed automatically): OD = 0.333 × stent_french wall_thickness = r_t × OD ID = OD − 2 × wall_thickness d_sh = min(r_sh × ID, ID − 0.15) section_length_mid = total − prox − dist
6/18

## Slide 7
Parametric CAD Generation
src/cad/stent_generator.py — scripted 3D geometry with build123d
What it does:
• Takes the 16 parameters as input
• Validates geometry constraints automatically
• Creates a proximal helix (pigtail coil)
• Creates a straight hollow tube body
• Creates a distal helix (second pigtail)
• Cuts side holes at computed positions
• Optionally cuts a half-pipe ("unroofed") section
• Exports to .STEP format → ready for COMSOL import
# Usage — it's this simple: from src.cad import StentGenerator, StentParameters params = StentParameters( stent_french=6.0, stent_length=150, n_prox=2, n_mid=5, n_dist=2, r_t=0.15, r_sh=0.5, unroofed_length=10.0 ) gen = StentGenerator(params) gen.export_step("my_stent.step") print(gen.get_info())
The generator handles all the math internally — helix sweeps, boolean cuts, hole positioning. You just give it the design and it spits out a 3D model.
7/18

## Slide 8
Latin Hypercube Sampling (LHS)
src/sampling/lhs_generator.py — covering the design space efficiently
Why LHS instead of random sampling? • Random sampling leaves gaps — you might cluster in one corner and miss another entirely • LHS divides each dimension into equal-probability bins and ensures exactly one sample per bin → guaranteed coverage • Think of it like a sudoku constraint applied to sampling: no two points share the same row or column in any dimension • We use scipy.stats.qmc.LatinHypercube under the hood • Default: 60 target samples × 3x oversample = 180 raw → filter down to ~60 valid designs
# The core of lhs_generator.py: from scipy.stats import qmc sampler = qmc.LatinHypercube(d=16, seed=42) samples_01 = sampler.random(n=180) # [0,1] # Scale to physical ranges scaled = qmc.scale(samples_01, lower_bounds, upper_bounds) # Round discrete vars (hole counts) for var in ['n_prox','n_mid','n_dist']: df[var] = df[var].round().astype(int)
The CLI script (run_lhs_batch.py) wraps this: python scripts/run_lhs_batch.py --n_samples 60 --seed 42 --output data/lhs_batches/batch_001.csv
8/18

## Slide 9
Feasibility Filtering
src/sampling/feasibility.py — rejecting impossible designs before simulation
Hard constraints (if violated → design is thrown out):
• Inner diameter ≥ 0.6 mm (too narrow = no flow)
• Middle section ≥ 10 mm (too short = pointless)
• Side holes must physically fit: n_holes × pitch ≤ usable length
(checked per section: proximal, middle, distal)
• Wall thickness minimum (structural integrity)
• Hole diameter can't exceed inner diameter
# How it works internally: class FeasibilityFilter: def _check_row(self, row): try: # Reuse StentParameters validation StentParameters(**filtered_args) return None # Valid! except ValueError as e: if "ID" in msg: return "ID < ID_MIN" elif "Cannot fit" in msg: return "Hole packing (mid)" ... # Returns a report: # FeasibilityReport( # n_input=180, n_valid=67, # rejection_reasons={ # 'Hole packing (mid)': 58, # 'ID < ID_MIN': 33, ...})
Why this matters: Each COMSOL sim takes 10–60 min. Filtering 180 → 60 saves potentially 100+ hours of wasted compute.
9/18

## Slide 10
COMSOL CFD Simulation
src/comsol/runner.py — automating the expensive part
What happens in this stage: 1. Take a .STEP file (from CAD generator) + design parameters 2. Import into COMSOL Multiphysics (industry-standard CFD solver) 3. Set up laminar flow physics — CSF (cerebrospinal fluid) flowing through the stent 4. Mesh the geometry, solve the Navier-Stokes equations 5. Export results: pressure drop (ΔP), total flow (Q_total), side-hole flows per section (Q_prox, Q_mid, Q_dist) The COMSOLRunner class automates this: • Constructs the batch command with parameter injection • Runs: comsol batch -input base.mph -output out.mph -pname "p1,p2" -pval "v1,v2" • Logs everything (convergence, CPU time, errors) • Can generate shell scripts for cluster deployment This is the bottleneck of the whole pipeline — each simulation takes real time. That's exactly why we need the GP surrogate: to AVOID running COMSOL for every design.
Think of COMSOL as the "ground truth oracle" — accurate but slow. The GP learns to approximate it cheaply.
10/18

## Slide 11
Parsing COMSOL Results
src/comsol/result_parser.py — extracting what we care about
@dataclass class COMSOLResult: run_id: str q_out: Optional[float] = None # Total outlet flow delta_p: Optional[float] = None # Pressure drop q_sh: Dict[str, float] = None # Side-hole flows converged: bool = False # Did it finish? iterations: int = 0 cpu_time_s: float = 0.0 mass_imbalance: Optional[float] # Sanity check errors: List[str] = None # What went wrong
What the parser does: 1. Reads the exported CSV from COMSOL (columns like Q_total, delta_P, Q_sh_prox…) 2. Scans the solver log for: • Convergence status ("Solver finished.") • CPU time ("Solution time: 123 s") • Iteration count • Error messages 3. Packages everything into a clean COMSOLResult object 4. Mass imbalance check: flow in ≈ flow out? (If not, something went wrong with the sim) This structured output feeds directly into the GP training data.
11/18

## Slide 12
What We're Measuring
The metrics from each simulation
Primary Metrics (optimization targets):
• ΔP — Pressure drop across the stent (minimize)
• Q_total — Total outlet drainage flow (maximize)
• Q_sh_prox/mid/dist — Flow through each section's holes
• CV_Qsh — How evenly flow is distributed (minimize)
Secondary / Diagnostic:
• WSS (Wall Shear Stress) — thrombosis risk
• Mass imbalance — conservation sanity check
• Solver iterations & CPU time — resource tracking
The Core Tradeoff (Pareto Front): You can't minimize pressure AND maximize flow at the same time — there's a tradeoff curve. Candidate Pareto fronts we'll explore: 1. ΔP vs Q_total (classic resistance vs throughput) 2. ΔP vs CV_Qsh (pressure vs flow uniformity) 3. 3D: ΔP vs Q_total vs CV_Qsh (the full picture) We start with single-objective optimization and graduate to multi-objective Pareto once the surrogate model is trustworthy.
12/18

## Slide 13
The GP Surrogate Model
src/surrogate/gp_model.py — learning to predict flow without simulating
What is a Gaussian Process (GP)? In plain English: it's a fancy curve-fitter that also tells you how confident it is in its predictions. • Feed it (design parameters, simulation results) pairs • It learns the relationship between them • For a NEW design, it predicts the flow metrics AND gives an uncertainty band ("I'm 95% sure ΔP is between 120–180 Pa for this design") Why not a neural network? • NNs need thousands of data points to train well • We have ~60–200 COMSOL results (each costs hours) • GPs are specifically designed for small-data regimes • The uncertainty estimates drive our optimization Kernel: Matérn 5/2 with ARD (Automatic Relevance Determination) → The model automatically figures out which parameters matter most
# Under the hood (gp_model.py): from botorch.models import SingleTaskGP from gpytorch.kernels import MaternKernel, ScaleKernel class GPModel: def fit(self, X_train, y_train): # Normalize inputs to [0,1] train_X_norm = self._normalize_x(train_X) # Standardize outputs (mean=0, std=1) train_Y_scaled = self._standardize_y(train_Y) # Matérn 5/2 kernel + ARD covar = ScaleKernel( MaternKernel(nu=2.5, ard_num_dims=16) ) self.model = SingleTaskGP( train_X_norm, train_Y_scaled, covar_module=covar ) # Optimize hyperparameters fit_gpytorch_mll(self.mll)
13/18

## Slide 14
Training & Validating the Surrogate
src/surrogate/training.py — making sure the model actually works
The GPTrainer handles the full workflow: 1. Train on available (X, y) data 2. Evaluate with multiple metrics: • RMSE — overall prediction error • MAE — average absolute error • R² — how much variance is explained • NLPD — uncertainty quality (penalizes overconfident wrong predictions) 3. Cross-validation (K-Fold): Split data into 5 chunks, train on 4, test on 1, rotate — gives honest estimate of generalization 4. Blocked holdout validation: Hold out entire batches (not random points) to respect the sequential nature of the data If the model isn't accurate enough → we go back and collect more simulation data (adaptive sampling)
# Training pipeline: trainer = GPTrainer() model, metrics = trainer.train(X, y) # metrics = { # 'train_rmse': 12.4, # 'train_mae': 8.7, # 'train_r2': 0.94, # 'train_nlpd': 2.1 # } # Cross-validation: cv_results = trainer.cross_validate(X, y, k_folds=5) # Returns per-fold RMSE, R², etc. # Evaluate uncertainty quality: def evaluate(model, X, y): mean, var = model.predict(X) # NLPD = 0.5*(log(2π·var) + (y-μ)²/var) # High NLPD = model is overconfident # and getting it wrong
14/18

## Slide 15
Literature Backing: Maulik & Taira (2020)
Phys. Rev. Fluids 5, 104401 — "Probabilistic Neural Networks for Fluid Flow"
What they showed: • Built "Probabilistic Neural Networks" (PNNs) that predict BOTH the flow value AND the uncertainty — just like our GP • Key finding: modeling fluid dynamics outputs as a Gaussian distribution works. The assumption holds. → This directly validates our GP approach. • Tested on 4 problems: 1. Shallow water equations (geophysical flow) 2. Cylinder vortex shedding (canonical CFD) 3. NACA airfoil wake (complex turbulence) 4. Sea surface temperature (real-world noisy data) • Their uncertainty estimates correctly flagged regions where the model was extrapolating → same principle drives our acquisition function
Why this matters for us: Taira uses a Neural Network to approximate a Gaussian process (outputs μ and σ). We use an actual Gaussian Process. Their work proves the Gaussian assumption is physically valid for fluid flow surrogates.
The data regime argument: Taira's PNN needs 100s–1000s of snapshots. We have ~60–200 COMSOL runs. GP is the mathematically correct choice for our "small data, high fidelity" problem. We're not disagreeing on physics — we're optimizing for our compute budget.
15/18

## Slide 16
Bayesian Optimization
src/optimization/optimizer.py — closing the loop intelligently
Once the GP surrogate is trained, we use it to decide WHERE to simulate next — not randomly, but strategically. The acquisition function (Expected Improvement / qEI): "Given what the GP knows so far, which untested design is most likely to be BETTER than the current best?" It balances: • Exploitation — designs the GP thinks are good • Exploration — designs the GP is unsure about (high uncertainty = worth checking) The optimizer: 1. Defines a scalarized objective (weighted Q_out + ΔP) 2. Optimizes qEI in the normalized [0,1] space 3. Denormalizes candidates back to physical units 4. Filters for feasibility (rejects impossible ones) 5. Returns the top candidates → simulate in COMSOL
# optimizer.py core logic: class BayesianOptimizer: def suggest(self, n_candidates=5): # Scalarized objective weights = {'q_out': 1.0, 'delta_p': -0.001} objective = ScalarizedPosteriorTransform( weights=weights_tensor) # Acquisition function acq = qExpectedImprovement( model=self.model.model, best_f=self._get_best_f(objective), posterior_transform=objective) # Optimize (oversample for feasibility) candidates, _ = optimize_acqf( acq, bounds=[[0]*16, [1]*16], q=n_candidates * 10, num_restarts=10, raw_samples=512) # Denormalize + filter valid = self.feasibility.filter( self._denormalize(candidates)) return valid[:n_candidates]
16/18

## Slide 17
The Full Campaign Script
scripts/run_optimization_campaign.py — tying it all together
This is the master orchestrator that runs the entire active learning loop: Step 1: Check for existing data → If none: generate initial LHS batch (with filtering) → If data exists: proceed to training Step 2: Train GP surrogate on all available data → Reports RMSE, R² to terminal Step 3: Run Bayesian Optimization → Suggests next batch of candidates Step 4: Generate CAD (.STEP) files for each candidate → Validates geometry, exports files Step 5: Save batch manifest CSV → Prints "NEXT STEPS: run COMSOL on these designs" Then you run COMSOL, append results, and run the script again → the loop continues.
# Campaign CLI: python scripts/run_optimization_campaign.py \ --campaign campaign_001 \ --batch_size 5 \ --init_lhs # (first time only) --n_init 20 # Output: # ========================================== # NEXT STEPS for Batch 0: # 1. Run COMSOL simulation for designs # listed in: # data/campaigns/campaign_001/batch_0000.csv # 2. Append results to: # data/campaigns/campaign_001/results.csv # ========================================== # After COMSOL, run again: python scripts/run_optimization_campaign.py \ --campaign campaign_001 \ --batch_size 5 # → trains GP on data, suggests NEXT 5
17/18

## Slide 18
Key Engineering Decisions
The "why" behind each choice
Objective
Start single-objective → graduate to Pareto once uncertainty is calibrated
Transforms
Inputs normalized [0,1]; log-transform outputs if they span orders of magnitude
Sampling
LHS 3× oversample → adaptive BO (space-filling baseline, then targeted)
Model
GP/Kriging over Neural Nets (works with <200 data points)
Kernel
Matérn 5/2 + ARD + noise (smooth-but-not-too-smooth; identifies key dims)
Training
Marginal likelihood (MLE) (stable, fast; upgrade to Bayesian if needed)
Validation
Blocked holdout (respects sequential batch structure)
Acquisition
EI/qEI → EHVI/qNEHVI (proven → Pareto-aware)
18/18

## Slide 19
What's Next
✅ Pipeline code is built and tested
✅ Parameter space defined (16 dimensions, config-driven)
✅ LHS sampling + feasibility filtering operational
✅ GP model wrapper + training pipeline ready
✅ Bayesian Optimizer implemented
⏳ Coordinate with Parnian on COMSOL model setup
⏳ Run first LHS batch through COMSOL (60 designs)
⏳ Train initial GP surrogate, validate on holdout
⏳ Begin active learning loop (BO → COMSOL → retrain)
⏳ Graduate to multi-objective Pareto optimization
⏳ Convergence → final optimal stent designs
19/18

## Slide 20
Questions?
Pipeline code: github.com/akashc/masters
Docs: docs/pipeline_overview.md | docs/decisions_gp_surrogate.md | docs/metrics_catalog.md
20/18